Material decomposition

Beer-Lambert equation

Ix,y,zo,E=Io(x,y,0,E)e-0zoμ(x,y,z,E)dz

where Ix,y,zo,E and Io(x,y,0,E) are intensities after and before traverse the sample, zo is the thickness of the sample, E is the energy of the X-ray beam, and μ is the attenuation coefficient. For a breast tissue with micro-calcifications, the integral term can be written as:

0zoμ(x,y,z,E)dz=μ1E0zog1x,y,zdz+μ2E0zog2x,y,zdz+μBE0zogBx,y,zdz

With μ1E, μ2E as the attenuation coefficients of two types of micro-calcifications, and μBE as the attenuation coefficient of the breast tissue. The energy-independent integrals, 0zogix,y,zdz, correspond to the projected thickness of each material. If 0zogTx,y,zdz=0zog1x,y,zdz+0zog2x,y,zdz+0zogBx,y,zdz is the total projected thickness of the sample, then the Beer-Lambert equation is

Ix,y,zo,E=Io(x,y,0,E)e-μB(E)0zogTx,y,zdz-∆μ1(E)0zog1x,y,zdz-∆μ2(E)0zog2x,y,zdz

fE=-lnIox,y,0,EIx,y,zo,Ee-μBE0zogTx,y,zdz=∆μ1E0zog1x,y,zdz+∆μ2(E)0zog2x,y,zdz

n images at different energies, fEi, are taken to retrieve 0zog1x,y,zdz and 0zog2x,y,zdz using matricial methods as:

fE1=∆μ1E10zog1x,y,zdz+∆μ2(E1)0zog2x,y,zdz

fE2=∆μ1E20zog1x,y,zdz+∆μ2(E2)0zog2x,y,zdz

⋮

fEn=∆μ1En0zog1x,y,zdz+∆μ2(En)0zog2x,y,zdz

∆μ1E1∆μ2(E1)⋮⋮∆μ1En∆μ2(En)0zog1x,y,zdz0zog2x,y,zdz=fE1⋮fEn→ΒΗ=Γ

And inverting the matrix equation for Η as

ΒTΒΗ=ΒTΓ

ΒTΒ-1ΒTΒΗ=ΒTΒ-1ΒTΓ→Η=ΒTΒ-1ΒTΓ

Verification

Images were acquired for two types of micro-calcifications—alumina and hydroxyapatite (HA)—at five different radii, embedded in a standard breast tissue with a thickness of 5 cm. The corresponding imaging energies were 20 keV, 24 keV and 28 keV (Figure 1), forming a 3×1 matrix for Γ. The objective is to retrieve the projected thickness of alumina and HA, using their attenuation coefficients at these energies to construct the matrix Β.

| E=20 keV | E=24 keV | E=28 keV |
| --- | --- | --- |

Figure 1.

| Alumina | HA |
| --- | --- |
| Theoretical |  |
| Material Decomposition |  |

Figure 2.

The results are presented in Figure 2, where the algorithm successfully retrieves the projected thickness of both materials. A comparison between the expected and retrieved projected thickness profiles—obtained by projecting the projected thickness images along the horizontal axis—is shown for both materials.

Characterization of the Timepix3 detector

ToT vs. E. Using Timepix3 in Pixel mode.

Figure 3. A: General diagram of dose measurement, illustrating the X-ray source and detector. B: General diagram used for characterization of the Timepix3 detector, to convert ToT to energy values. The detector is arranged so that a metal can read the fluorescence photons generated by it. The ToT distribution obtained for comparison.

Use 4 metals: Emission lines sufficiently separated, such that they cover a wide range of energies.

- Zr →Eα1=15.775 keV →V=30 kVp

- Rh →Eα1=20.216 keV  →V=50 kVp

- Sn →Eα1=25.271 keV →V=50 kVp

- Gd →Eα1=42.996 keV →V=80 kVp

Figure 4. Left. ToT spectrum of four materials. Right. Dependence of ToT with energy.

Spectrum of the X-ray source using the Timepix3.

ME=U(E,E')S(E')dE'

- ME: Energy spectrum measured by the detector

- SE: Spectrum of the X-ray source

- U(E,E'): Detector response function

Obtaining U(E,E'), is determined by the energy resolution of the detector. This resolution is determined by the fluorescence spectra of four materials: Zr (Eα1=15.775 keV), Sn (Eα1=25.271 keV), Gd (Eα1=42.996 keV), Am241 (Eγ=59.541 keV).

- Perform a Gaussian fit around the peaks of the spectrum

- Obtain σ and Eo from each fit to obtain the energy resolution from each material. →RE=FWHMEo=22ln2σEo

- Graph of RE vs E

Figure 5. Left. Energy spectrum of four materials, with Gaussian fits. Right. Energy resolution.

From the energy resolution, it is possible to construct each element of U(E,E'), as illustrated in Figure 6.

Figure 6. Left. U(E,E') for E=8 keV (Black), E=17 keV (Blue) and E=27 keV (Red).

Consequently, SE is obtained for each given spectrum using the Bayesian deconvolution method, as illustrated in Figure 7 for Rh and Al at 28 kVp.

Figure 7. Top. ME measured for Rh and Al at 28 kVp. Bottom, SE retrieved for Rh and Al at 28 kVp.